## Slide 1
Large Frame Transfer Detectors for the MAIA imager
Jesús Pérez Padilla1,2, Simon Tulloch3, Gert Raskin2, Saskia Prins1,2, Florian Merges1,2
1 Mercator Telescope, Roque de los Muchachos observatory, La Palma, Canary Islands, Spain 2 Instituut voor Sterrenkunde, KU Leuven, 3001 Leuven, Belgium 3 QUCAM, Cádiz, Spain
Mercator Telescope
Institute of Astronomy
 | Slow (LGN) | Fast (MGN)
Read-out frequency | 156kPixel s-1 | 227kPixel s-1
Read-out noise | 3.5 - 4 e- | 4.5 - 5 e-
Single row read-out time | 13.7 ms | 9.5 ms
Full frame read-out time | 43.1 s | 29.8 s
Full frame transfer time | 295 ms | 295 ms
Conversion gain | 0.8 ADU/e- | 0.4 ADU/e-
MAIA
Mercator Adv. Instr. for Ateroseismology
3-channel (SDSS u’, g’, r’ bands).
Frame transfer detectors
2048x6144 13.5µm pixels. 2k x 3k imaging area, and 2k x 3k storage area.
Read-out characteristics
System uses the FT characteristic to allow different exposure times for the three channels.
Linearity study and optimization
Detectors showed a non-linear behavior at lower levels of signal. Problem was solved by controller timing code optimization, aided by an accurate method for photometric analysis.
Acquisition system
ARC GEN-III controller. Two ARC-45 Video boards, and two ARC-32 Clock Driver boards
Gain error ≈ ±5%
R
U
G
Linearity down to very low signal levels (<10 ADU)