## Slide 1: Characterisation of CCDs for the EUCLID VIS channel Peter Verhoeve, Thibaut Prod’homme, Nathalie Boudin Payload Technology Validation Section Future Missions Preparation Office Directorate of Science and Robotic Exploration of the European Space Agency

## Slide 2
OUTLINE
The EUCLID mission and our role
EUCLID VIS CCD description
Pre-irradiation Characterisation
Irradiation and post-irradiation characterisation
Tests to come
Summary

## Slide 3
Euclid (ref. [1]) : ESA mission to map the geometry of the dark Universe
Using two cosmological probes, namely Weak Lensing and Baryonic Acoustic Oscillations
1.2m diameter telescope with two focal plane instruments provided by the Euclid Mission Consortium:
Near-infrared spectrometer-photometer (NISP) (ref. [2]),
Visual imager (VIS) (ref. [3]) :
  will be used to measure the shapes of galaxies in one single wide visual band (550-920 nm)
  focal plane array of 36 back-illuminated CCDs (4k×4k pixels each) with 0.101 arcsec pixel platescale, giving a geometric field of 0.55 deg2
With the weak lensing technique, the mass distribution of the lensing structures can be traced back
Originally baselined CCDs were e2v CCD203-82.
Following the results from a dedicated CCD devices.
ESA survey mission to map the geometry of the
dark Universe
Using two cosmological probes:
Weak Lensing + Baryonic Acoustic Oscillations
launch 2020, L2 orbit, 6yr nominal mission
1.2m diameter telescope with two focal plane instruments provided by the Euclid Mission Consortium:
  Near-infrared spectrometer-photometer (NISP),
  Visual imager (VIS):
    will be used to measure the shapes of galaxies in one single wide visual band (550-920 nm)
    focal plane array of 36 back-illuminated CCDs (4k×4k pixels each) with 0.1 arcsec pixel plate scale, for a geometric field of 0.55 deg2
With the weak lensing technique, the mass distribution of the lensing structures can be traced back
Euclid requires the accuracy with which the shape of the galaxies is to be measured is 1%.
The radiation damage effects will compromise this accuracy, and thus need to be characterized.
EUCLID

## Slide 4
The EUCLID VIS intrument
Courtesy EUCLID Consortium/VIS

## Slide 5: Motivation
ESA has run a contract for the pre-development of the CCDs during the EUCLID definition phase and ESA is responsible for the procurement of the flight CCDs for the EUCLID VIS instrument
A CCD test bench was developed in the Payload Technology Validation section (old SRE-FI) during the pre-development phase in support of the definition phase
The ESA Euclid project team uses now the now fully operational facility to:
    - verify other test results produced in the flight production phase - support the EUCLID VIS consortium in case of issues.
This work:
  Characterise a representative CCD, pre- and post- proton-irradiation
  In particular, try to quantify the change in the measured shape of galaxies due to radiation damage
  Compare with/feed into simulations and corrective models

## Slide 6
CCD description: CCD273
The CCD273 is developed for EUCLID-VIS based on existing CCD203 (e2v Technologies)
4kx4k format (5x5 cm2)
12x12 micron pixel size
4 high-responsivity, low-noise read-out circuits
high-res silicon, back-thinned to 40 micron
Thin gate dielectric process
Image section split in two, with charge injection structure in the middle
Reduced register width for rad-hardness
AR coated
SiC package
Specification:
Amp responsivity 6-8 µV/e-
Read noise <3.6 e-rms @70 kHz
FWC >175ke-
CTI <5e-6
Dark signal <6e-4 e-/s/pixel @ 153K

## Slide 7
Noise and conversion gain
DN to electron conversion from 55Fe spectra (Mn-Ka 5.9 keV yields 1580 e- @153K)
CDS (Dual Slope Integration), 3.4 µs integration time
70 kHz pixel rate (14.3 µs pixel time)

## Slide 8
Noise and conversion gain
Noise 2.5 e- rms @ Vod =27.0 V, 70 kHz pixel rate
For all 4 amplifiers

## Slide 9
Channel parameter
~1.7 V difference in channel parameter between left and right EH: 10.6V
FG: 8.9V
Resistivity = 0.90
The cause of the variation of the channel potential has been traced and eliminated @e2v

## Slide 10
Dark current
Measured with:
  - 10x line binning
  - 0-3000s integration time
Thermal down to 170K
<10-4 e-/s/pixel @ 153K

## Slide 11: Set-up for monochromatic illumination (uniform or spot)
RD1 movable
Abs. int. mapping
Integrating
Sphere
Oriel 70401
RD2 fixed
Rel. int. monitor
Light tight enclosure
CCD
Dewar
Light source
(Xe, 100 W)
Program-
mable
shutter
grating
Monochromator
=200-1060nm
Adjust.
Slit =
2-20nm
Straylight
screen
AR coated
Fused silica
window
Source
aperture
Photodiode
aperture
CCD

## Slide 12
Quantum Efficiency
Measured QE ~3% above e2v measurements
Estimated uncertainty ~3%

## Slide 13
Charge transfer inefficiency (CTI)
from x-ray illumination (55Fe, 5.9 keV):
CTI at 153K
@ T=153K, 80 pixels/photon:
serial CTI =4.0ˑ10-6 parallel CTI =~1.0ˑ10-6

## Slide 14
Charge transfer inefficiency (CTI)
from x-ray illumination (55Fe, 5.9 keV):
CTI vs temperature
@ 80 pixels/photon:
serial CTI: better at lower T parallel CTI: no measurable T dependence

## Slide 15
Proton Irradiation at the KVI accelerator
Groningen, Netherlands
CCD at ambient temperature, unbiased
10 MeV protons, flux 9.0 106 p+ cm-2 s-1
1st illumination: Region I: 2.40 109 p+ cm-2.
2nd illumination: Region II+III: 4.80 109 p+ cm-2 (= predicted EUCLID EOL dose)
Dark image at 243K

## Slide 16
Dark signal, post-irradiation

## Slide 17
Charge transfer inefficiency (CTI)
from x-ray illumination (55Fe, 5.9 keV)
Post-irradiation
Cf N. Murray, Proc. of SPIE Vol. 8453 845317-2
Optimum temperature ~150K
(EUCLID ~153K)

## Slide 18
Charge transfer inefficiency (CTI)
Radiation Damage constants
Radiation damage constants
(@ 153K, 70 kHz, ~70 pixels/5.9keV photon)
ΔCTI/(10 MeV proton fluence) =
Serial: 5.4e-15
Parallel: 9.8e-15
Cf CEI, OU (FI, @160K, 200 kHz):
Serial: 3.9e-15
Parallel: 11.0e-15

## Slide 19: Trap Pumping
Flat field illumination ~4500 e-
Followed by ~4000 pairs of forward and backward line shifts reveals traps
Proposed as in orbit calibration tool for EUCLID-VIS
(N. Murray, Proc. of SPIE Vol. 8453 845317-2)

## Slide 20
CTI from Extended Pixel Edge Response
(EPER) at uniform illumination
EPER data to be used for comparison with charge loss models
T

## Slide 21: Next step: the ultimate experiment for EUCLID
Project a sky scene of stars and Galaxies on different sections of the CCD
Measure the shape of the galaxies in EUCLID observables
Determine the change in shape for different levels of irradiation

## Slide 22: Next step: the ultimate experiment
Lithographic mask
Bright stars + galaxies
Galaxies in different shapes, sizes and orientations
~400 galaxies and 24 stars
Average distance between objects ~100 pixels
Mask just arrived
Testing in next few months
5x5mm

## Slide 23: Summary
We have characterised a EUCLID prototype back-illuminated CCD273
We have irradiated this CCD with 10 MeV protons up to the expected End Of Life Dose of 4.8e9 protons/cm2
The observed increase in Charge Transfer Inefficiency (5.9 keV photons, ~60 pixels/photon) corresponds to:
  1.0e-5 per 1e9 protons/cm2 for parallel transfer
  0.5e-5 per 1e9 protons/cm2 for serial transfer
  In good agreement with results on a front-illuminated device (Centre for Electronic Imaging, Open University, UK)
Next step is to measure the radiation induced changes in the shape of projected elliptical galaxies (in EUCLID observables)