## Slide 1: Characterization of a Large Format HgCdTe on Silicon Focal Plane Array
B. Hanold, J. Lee, D. Figer – Rochester Institute of Technology
L. Mears, J. Bangs, E. Corrales, J. Getty, C. Keasler, M. Mitani – Raytheon Vision Systems

## Slide 2: Outline
Project overview
Device design
Test setup
Characterization results
Going forward

## Slide 3: Project Overview
HgCdTe detector cost can be reduced by using Si substrate instead of CZT
Project will develop low cost infrared detectors for astronomy with long term goal of producing larger arrays enabled by using larger Si wafer
Project goal is to fabricate 2K x 2K MBE HgCdTe/Si detectors with competitive performance
Work is being funded by NSF and NASA to develop large format HgCdTe/Si detectors in collaboration with Raytheon Vision Systems (RVS)

## Slide 4: Device Design Drivers
Dark current and quantum efficiency identified as drivers for HgCdTe/Si design improvements
Multiple pixel designs need to be tested
Large amount of testing required to select optimal design

## Slide 5: Variable Unit Cell (VUC) Devices
4 1K x 1K variable unit cell detector die
1K x 1K die fabricated with 4 unit cell designs
Design allows direct comparison of detector characteristics
VUC detector speeds design selection and allows more time for optimization

## Slide 6: Detectors To Date
CfD received 4 detectors from RVS:
  SN: 9A, 14, V1, and V2
All bonded to Virgo ROICs
Current progress:
  Characterization of 2K x 2K HgCdTe/Si (SN: 14)
  Characterization of 1K x 1K HgCdTe/Si (SN: V1)

## Slide 7: Rochester Imaging Detector Laboratory (RIDL) in The Center for Detectors (CfD)
3 cryogenic test systems
Computing cluster for data acquisition and reduction running automated test suite
Test systems integrate with telescope for on-sky evaluation of detectors
Test systems have been designed and used for radiation testing
Test Results for an Array-Based GM_APD Detector Before and After Irradiation
K. Kolb’s Poster L10

## Slide 8: Cryostat Detail
Detector Enclosure
Cold Plate
Filter Wheel
Outer Case
Electrical Connectors

## Slide 9: Detector Control Electronics
Potentiometer
SMT JFET
Through Hole JFET
SMT resistor
ARC Gen III controller used to operate the detectors
Mezzanine current source board designed for output buffer current supply
~7e- CDS noise - includes cabling
Noise not increased significantly with addition of current source circuit

## Slide 10: Virgo-14 Read Noise
18 e- read noise CDS
5.5 e- read noise Fowler-16
Noise may be improved with bias noise reduction

## Slide 11: Virgo-14 Well Depth, Non-linearity, and Gain
Well depth: 126 K e-
Non-linearity terms:
  a = 1.712E-6
  b = -1.59E-11

## Slide 12: Virgo-14 Crosstalk
Crosstalk measured using cosmic rays in dark exposures
3 x 3 grid shows crosstalk in nearest neighbors around central hit
Results given in percentage of hit signal
H2RG-015-5.0µm
SB304-008-5.0µm
VIRGO-14-4.9µm
Asymmetric crosstalk due to incomplete settling

## Slide 13: Virgo-14 Dark Current
Virgo-14 produced for a previous RVS project
Device being used as benchmark to compare future devices against
Currently measuring QE and validating 4.9 µm cut-off
.02 e-/s/pixel

## Slide 14: Going Forward
Project Milestones:
  Characterize VUC detectors
  Characterize substrate removed VUC detectors
  Select pixel design using VUC detector performance
  Fabricate and characterize 2K x 2K substrate removed MBE HgCdTe/Si device
Long term goals for MBE HgCdTe/Si process:
  Scale design to 4K x 4K and 8K x 8K
  Reduce pixel pitch

## Slide 15: CfD Personnel
P.I.
Don Figer
BS, ME
Kenneth Bean
Brandon Hanold
Engineer
Joong Lee
Engineer
Engineer
Iain Marcuson
Matt Davis
MS, EE
PhD, IS
Kim Kolb
Mike Shaw
BS, EET
BS, EE
Zach Mink
BS, Physics
Mike Every
MS, EE
Jon Zimmermann
Notes: 15

## Slide 16: Thank you for your attention.
Questions?